Technology parameters of LP-1900 infusion pump

| Pumping way | Semi-squeeze finger-type peristaltic pump |
| --- | --- |
| Infusion sets | Special or quality PVC/TPE infusion sets with an external diameter of 3.4mm~4.5mm |
| 4 infusion modes | Flow-rate, Time-amount, Drop-rate, Body-weight |
| Flow-rate mode | Flow-rate: 1.0～1200.0 ml/h，0.1ml/h step | Preset volume: 0.1～9999.9ml，0.1ml step Cumulant: 0～9999.9ml，reset only |
| Time-amount mode | Time: 1s~99h59m59s,1s step |  |
| Drop-rate mode | Drop-rate: 1~333 d/min, 1 d/min step |  |
| Body-weight mode | Dose: 0.1～999.9 mg，0.1mg step |  |
|  | mg/kg/h: 0.01～999.99 mg/kg/h，0.01 mg/kg/h step |  |
|  | Body-weight: 1～300 kg，1kg step |  |
| KVO function | KVO rate: 1.0～5.0 ml/h，0.1 ml/h step |
| Bolus funtion | Bolus rate: 50.0～800.0 ml/h，0.1 ml/h step |
|  | Bolus volume：0.1～99.9 ml/h，0.1ml step |
| Infusion precision | Better than ±5% (using a calibrated or a special infusion set) |
| Alarm | 1,Air in tube, 3 grades of sensitivity; 2,Bubble sensor fault; 3.1,Upper-end occlusion, -20～-60 kPa,8 grades of sensitivity; 3.2,Upper-end base-pressure fault; 4.1,Lower-end occlusion, 20～140 kPa,8 grades of sensitivity; 4.2,Lower-end base-pressure fault; 5.1,Pump rotation fault; 5.2,Inverse-infusion; 6,Abnormal door-opening; 7,Low battery power, battery fault; 8,Overtime pause; 9,Completed infusion; 10,AC power fault; 11,MCU fault, and stop infusion. |
| Data function | 1,Memory of calibration：Records the calibration data of tube automatically for application; 2,Memory of parameter：Records the latest infusion parameter for next operation; 3,Historical records: records latest 1000 infusion events. (what/when/why,flow-rate, etc.) |
| Safety classification | Class I; Internal power supply, CF type, IPX4 equipment |
| Power supply | AC 100~240V, 50/60Hz, Built-in Ni-Mh rechargeable battery |
|  | AC power consumption: ≤20VA |
| Weight and size | 130mm(w)×206mm(h)×118mm(t);≤1.6kg |